## Sheet1
PRODUCT产品 | TYPE型号 | NEEDED需要数量 | TECHNICAL SPECIFICATIONS技术规格 | 建议报价 CIF 罗马尼亚主要空港
BODY SUIT 紧身衣 | ATI | 3000000 | minimum class 4 protection against viruses; waterproof at > 2 bar presuure; minimmum class 3 according ISO 16603 最低4级病毒防护；压力大于2巴时防水；最低3级符合ISO 16603 | USD 22.50
FFP3/N99 | medical医用的 | 10000000 | medical grade ffp3. CE or FDA; protection minimum 99% of particles < 0,5 microns according EN 149/2001 or class FFP3/P3, NIOSH - N99医学级ffp3。CE或FDA；根据EN 149/2001或FFP3/P3级、NIOSH-N99，99%的保护颗粒<0.5微米 | USD 3.6
FFP2/N95 | medical | 10000000 | medical grade ffp2. CE or FDA; EN 149:2001+ A1:2009 医学级ffp2。CE或FDA；EN 149:2001+A1:2009 | USD2.10
Protective Glasses 防护眼镜 | Medical | 200000 |  | USD6.15
Protective face shield / Visors防护面罩/护目镜 |  | 200000 | CPV code: 35113410-6 - Protective clothing against biological and chemical agents, comply with SR EN 166, Disposable protective visor 4.1.1 Have an adjustable / elastic head restraint system, which ensures easy visor placement and removal and ensures that the visor is firmly held in the figure; 4.1.2 The visor material shall be transparent and resistant to splashes and liquid drops, as well as to medium energy impact, in accordance with SR EN 166 (including in the form of aerosols). | USD7.32
Single-use surgical apron一次性手术围裙 |  | 2000000 |  | USD2.50
NOVA-LUNG iLA (gas exchanger for arteriovenous carbon dioxide removal) and compatible kitNOVA-LUNG iLA（动静脉二氧化碳清除用气体交换器）和兼容试剂盒 |  | 20
Monitors for complex vital functions复杂重要功能监视器 |  | 200
Video system for difficult intubation and bronchoscopy困难插管和支气管镜检查的视频系统 |  | 200 |  | CIP USD 18800
Videolaringoscop视像仪 |  | 5000 |  | CIP USD 12700
Infusomat漏斗 |  | 1000 |  | CIP USD 1860
Simple ventilation CPAP简易通气CPAP |  | 1000 |  | CIP USD 1780
Invasive ventilation有创机械通气 |  | 1000 | "Ventilation mode: A/C, SIGH, SIMV(SIMV+PSV), SPONT+PSV Tidal volume: 0 1500mL Breath rate: 4 100bpm SIMV rate: 1 40bpm Inspiration time: 0.1 12s Pause time: 0 4s Trigger sensitibity: -20 0cmH2O PSV: 0 60cmH2O Pressure range: 5 60cmH2O High pressure: 20 80cmH2O Low pressure: 0 20cmH2O Ventilation waveform: Square and decelerate waveform PEEP: 0 40cmH2O FiO2: 21% 100% Alarm silence: ≤120s Parameters monitored: Vti, Vte, MV, FiO2, Ppeak, Pmean, PEEP, Ftot, Fspont, MVspont Waveforms displayed: P-T, F-T Alarm: O2 failure, air failure, MV high/low limit, Vt high limit, Fspont high limit, Paw high/low limit, apnea, FiO2 high/low limit, power failure Max. alarm record: 100 Safety valve: ≤8KPa Patient type: Adult & pediatric Driven mode: Pneumatically driven and electronically controlled Setting: Navigator wheel knob & touch key Display: 10.4''TFT color screen Power supply: 220V, 50Hz 110V, 60Hz Power consumption: ≤80W Gas supply: O2, 0.3 0.5MPa, 50L/min Dimensions(LWH): 350mm400mm300mm Weight: 15kg" | CIP USD 19800